DC 44

ELECTRICAL DIALYSIS AND CHEMOTHERAPY CHAIR, 4 MOTORS

- 4 motors for electronic adjustments of backrest, legrest, height, and Trendelenburg functions by hand control unit

- Suitable for Hemodialysis and Chemotherapy sessions

- Anti-microbial, water proof, flame retardant, air permeable and easy cleanable cushion and not effected to skin

- The chair is upholstered by viscoelastic foam which is known as memory foam, is a type of open cell, flexible polyurethane foam. It is body heat and weight sensitive material. It contours to the body shape, distributes the body pressure over a wide area and besides comfortable seating, allows for good blood circulation

- Foam’s capacity let the patient redistribute weight and surface pressure, potentially reducing decubitus ulcers

- ABS covering designed to keep itself in the upholstery and it blocks the dust which goes into the upholstery

- During the back rest movement, arm rests moves with the back rest simultaneously and the arm rests are easy removable

- Easy adjustable and removable upholstered arm rests

- Adjustable and removable cramp support and head support

- Head support provided with cushion for neck support

- 360° swivel, diagonal lockable castors of 125 mm diameter

- Height adjustable IV pole with 2 hooks

- Food / Reading Table (optional)

- Rechargeable Back-up battery (optional)

- Weighing Scale (optional)

- Central brake castors, 125 mm (optional)

- Manual CPR (optional)

- Tablet with holder (optional)

- Reading lamp (optional)

|  | TECHNICAL DETAILS |  |
| --- | --- | --- |
|  | Backrest Angle | : 0°-70° |
|  | Trendelenburg Angle | : 0°-13° |
|  | Reverse Trendelenburg Angle | : 0°-(-5)° |
|  | Legrest Angle | : 0°-35° |
|  | Safe Working Load Capacity | : 180 kg |
|  | Weight | : 115 kg |
|  | External Dimensions (with Armrest) | : 201 x 105 cm |
|  | Bed Surface | : 201 x 65 cm |
|  | Height Range | : 55-75 cm |
|  | Castor Diameter | : 125 mm |

|  | ELECTRICAL DATA |  |
| --- | --- | --- |
|  | Power In | : Max. 3,15A 100V-240V ~ 50-60 Hz |
|  | Electrical Shock Protection | : Class I, Type B |
|  | Liquid Ingress Protection | : IP66 |

ACCESSORIES

Rechargeable Battery Backup

ABS Food/Reading Table

HPL Food/Reading Table

Weighing Scale

Tablet with Holder

Central Brake Castors 125 mm

Reading Lamp

Pushing Handle

Manual CPR

COLOR CHART